Calibration of the Ocular Micrometer on a Microscope

Microscopes are often used to measure small objects.  For instance, forensic scientists use microscopy to measure the distance between microscopic tool marks or the diameter of fibers.  As you can see in figure 1 below, compound microscopes have ocular micrometers, or rulers that can help you measure items under the microscope.  However, the scale on the ocular micrometer changes with total magnification, and thus has no absolute value. Therefore, the ocular micrometer does not have units and it needs to be calibrated prior to use.  In this activity, you will learn how to calibrate the ocular micrometer, a skill that will help you in next week’s microscopy lab.

Figure 1:  Ocular micrometer superimposed over a view of red blood cells under the microscope.

You will use a stage micrometer to calibrate the ocular micrometer.  A stage micrometer is essentially a ruler that is mounted on a microscope slide that does have units (millimeters (mm) or micrometers (m)).  When calibrating, you will line up the stage micrometer with the ocular micrometer and count the number of divisions on the ocular micrometer per millimeter or micrometer on the staged micrometer.  The number of divisions will change as the magnification changes.

Figure 2:  Comparison of the ocular micrometer and the stage micrometer.

Example:  At a total magnification of 40x, a student measured 42 ocular micrometer divisions per millimeter.  What is the distance in micrometers per ocular unit?

1mm42 ocular unitsx1000μm1mm=23.8 m per ocular unit

Problem 1:  At a magnification of 40x, a student measured 41 ocular micrometer divisions per millimeter.  What is the distance, in micrometers, per ocular unit at 40x?

Problem 2:  At a total magnification of 100x, a student measures 16.4 ocular micrometer divisions per millimeter.  What is the distance, in micrometers, per ocular unit at 100x?

Problem 3:  At a total magnification of 400x, a student measures 4.1 ocular micrometer divisions per millimeter.  What is the distance, in micrometers, per ocular unit at 400x?

Problem 4:  Notice the pattern in the above 3 problems.  What do you think the distance in micrometers per ocular unit at 1000x would be?

Problem 5:  It was found that the units on the ocular micrometer and the units on the stage micrometer matched up 40 units to 1.0 mm at 100 power.  (So 40 ocular units equals 1.0 mm.) If the wing of a dead fruit fly was measured to be 2.5 ocular micrometer units at a magnification of 100, what is the length of the wing in mm? What does mm stand for?